## Sheet1
Mass Spec Operating Procedure
Do not run fractions on the mass spec until you have ran a gel. The mass spec is for final confirmation of purified protein.
Sample Preparation
1. Make 1 uM sample by diluting sample in buffer A/water.
2. Centrifuge for 1 min at high rpm to remove bubbles/ remove any precipiate.
3. Transfer to sample vial, ensure no bubbles are at the bottom of the vial. Only use the vials listed below ( Some sample vials are not letting the needle through)
o Thermo Scientific™ National™ Limited-Volume Polypropylene Vials with 10-425 Finish, Cat no:03-375-17CA, No.: C401011
o CAP W/SEPT BLU PTFE/SIL/TFC/PK Cat no:03-375-29A. No.:C4010-40A
Set-up
1. Put samples in autosampler
2. Get standard from 4 degree fridge in 1019. (can skip calibrating if already used that day)
3. Fill Syringe with standard and reattach.
4. Go to oftcontrol and turn system from standby to operate
5. Hold in the forward arrow and the run button at the same time for approximately 5 seconds. You should see the standard masses show up on the computer screen.
6. Hit the run button with the standard and make sure the arrow is flashing on the screen, this shows the standard is running.
7. Click box to turn UV light on (it takes a few minutes to fully turn on)
8. Go to Compass Hystar and click on sample table.
9. Fill out the table with the vial locations, sample name, folder and program.
10. Click save as: fill in the tag and name
11. Once saved click "Que for Acquisition" in the bottom right corner
12. The sample table will be on the left side under the sample table que. Double click on the sample table to bring it up on the middle screen.
13. Go to the oftcontrol and near the bottom left click calibrate, if the percentage is green click accept.
If the percentage is a different color click on the mass with the highest error and you can clear it from the calibration. you can do this until the percentage is green. Then accept
14. Go back to the hystar screen and click start.
15. Once the run starts fill in the logbook.
Trouble Shooting
1. Make sure your sample was taken (if it wasn't, let Christy know and skip other trouble shooting steps)
2. Try a 2 uM sample if you do not see your sample at 1 uM.
If you do not let Christy know there is an issue, the mass spec won't be tended to.